# PERFORMANCE REQUIREMENTS FOR 
HydroLOGical / Meteorological EQUIPMENT IN VTS

### INTRODUCTION

In 1997 the IMO Maritime Safety Committee adopted Regulations for Vessel Traffic Services (VTS) that have since been included in SOLAS Chapter V (Safety of Navigation) as Regulation 8-2.This Regulation specifies the responsibilities of contracting governments to arrange for the establishment of VTS in certain vulnerable areas under their control.

### BACKGROUND

The environmental systems in a VTS, also referred to as the hydrological / meteorological (hydro/meteo) systems, include sensors and readouts of various meterological or hydrographical variables.  Typical meteo variables are those provided by weather stations and include air temperature and humidity, wind velocity and direction, and visibility.   In certain locations, hydro variables such as tidal level, and current direction and velocity may be required.  These data may be obtained through sensors or available in tables/databases from national authorities.  Sensors providing this data, usually located at remote sites, communicate the variables to a Vessel Traffic Centre (VTC) via a telecommunications link.  At the VTC, graphical and/or numeric information is presented for use by the operators.

### OBJECTIVES OF HydroLOGical / Meteorological EQUIPMENT

Digital hydrological and meteorological information may be integrated into VTS applications to provide the operator a real-time assessment of the environmental situation in the VTS area of responsibility.  Information collected from this equipment can be provided to ships to assist in assessing the waterway conditions.

### RELATED REQUIREMENTS INVOLVED

In this section all additional applicable IMO requirements for the respective navigational equipment should be listed. This may be:

Relevant SOLAS requirements; SOLAS Chapter V (Safety of Navigation) as Regulation 8-2

General resolutions, such as resolution A.694(17) - General Requirements for Shipborne Radio Equipment forming Part of the Global Maritime Distress and Safety System (GMDSS) and for Electronic Navigational Aids;

Resolutions with more specific requirements, such as:

Resolution A.688(17) - Code on Alarms and Indicators;

Resolution MSC.39(63) - Adoption of amendments to the Code on Alarms and Indicators.

World Meteorological Standards (TBD)

### DEFINITIONS

If necessary in this section terms used in the individual performance standards should be defined.

### FUNCTIONAL REQUIREMENTS

6.1	Sensors and their availability

### The measurement sensors must be installed and located according to recommendations from a VTS authority in consultation with hydrologist/meteorologist(s).

The most common measurements are: 
wind speed
wind direction

Wind gust
air temperature
relative humidity
precipitation

air pressure

visibility
water temperature
seawater level

height of tide
current (may be required at various depths)

wave height

wave direction

6.2	Reliability, accuracy and discrimination

The accuracy of the measurements must satisfy the following minimum requirements as determined by WMO:

Wind speed		+/- 0,5 m/s over the full range of 0-200 mph (???m/s)

Wind direction 		better than +/- 5°(2°??) over the full range of 0°-360°

Wind gust

Air temperature		resolution better than 0,1°C over the full range of -30°C to +50°C

Air pressure		+/- 0,2 hPa

Visibility		From 33 ft to 19 miles (10 m to 30 km), accurate to +/- 5%

Air humidity		+/- 3% over the full range of 0-99%

Precipitation		+/- 0.5%

Water temperature	better than 0,1°C

Tidal Level		0-29 meters

Current			+/- 5° over speeds from 0-350 cm/sec (0-8 ??? kts), accurate to 0.2 cm/s (???kts)direction

Wave height

Wave direction

6.3	Malfunctions, warnings, alarms and indications

In this section reference should be made to relevant requirements of resolution A.686(17) and resolution MSC.39(63). Depending on the individual type of the navigational equipment requirements in addition to resolutions A.686(17) and MSC.39(63) should be determined if necessary.

6.4	Software requirements

Not applicable

### OPERATIONAL REQUIREMENTS

7.1 Ergonomy

Not applicable

7.2	Operational controls

Not applicable.

7.3	Presentation of information

The results of the measurements must be presented in approved (SI) units.

The measurements must be available to the traffic operators at all times through an integrated display or separate instruments.  Data may be presented numerically and/or graphically.  A log of the latest 24 hour measurements must be available to the operators either numerically or graphically.

7.4	Software requirements

Not applicable

### DESIGN AND INSTALLATION

8.1	Durability and resistance to environmental conditions

External installation of electronics must be installed in an environmental enclosure .  As regards environmental conditions the IEC requirements' should be applied as far as relevant.

8.2 	Interference

These sensors are industrial equipment and therefore comply with applicable international standards and regulations.  In this section reference should be made to the IEC requirements (IEC 60945?) as far as relevant.

8.3	Power supply

IEC requirements should be applied as far as relevant.  In remote locations, due to the low power consumption of hydro/meteo sensors, authorities should consider use of alternative power (e.g., solar panels, wind vanes, etc.) in lieu of generators when commercial power is not available.

8.4	 Installation

Requirements concerning the installation of sensors, wiring and the arrangement of the equipment providing hydro/meteo information to the VTC should be determined.  Operational requirements will determine where sensors are to be located and how many are required.  Sites for sensors should consider the optimization of data to the VTS, as well as maintenance and repair.  Other considerations include availability of power, protection against vandalism, housing, and collocation with existing VTS, AtoN, or other suitable infrastructure

'IEC 529 "Degrees of protection provided by enclosures (IP Code)"

IEC 721-3-6 "Classification of environmental conditions, Part 3: Classification of groups of environmental parameters and their severities; Ship environment"

IEC 60945 "Maritime navigation and radiocommunication equipment and systems - General requirements, methods of testing and required test results"

8.5 Maintenance

Requirements in addition to resolution A.694(17) concerning maintenance should be determined if necessary. Sites for sensors should consider the maintenance and repair requirements, including remoteness and accessibility.

### INTERFACING

Requirements concerning the correctness and unambiguity of input and output data of the individual components should be determined. Relevant existing standards' should be taken into account.

### BACK-UP AND FALL-BACK ARRANGEMENTS

Depending on the individual type of the equipment, requirements concerning back-up and fall-back arrangements should be determined if necessary.

### SAFETY PRECAUTIONS

Depending on the individual type of the equipment, requirements in addition to resolution A.694(17) concerning safety precautions should be determined if necessary.

### MARKING AND IDENTIFICATION

Depending on the individual type of the equipment, requirements in addition to resolution A.694(17) should be determined if necessary, e.g., for the identification of software.

### DOCUMENTATION

In this section the contents of the documentation (e.g. Operator's Manual, Installation Manual, Service Manual) necessary for the operation of the equipment should be determined. The documentation should contain the following items:

General:

manufacturer, type of equipment and the serial number; system description;

specification of the equipment; supplementary equipment;

National Marine Electronics Association (NMEA)

Radio Technical Commission for Maritime Services (RTCM)

American National Standard Institute (ANSI)

Electronic Industries Association (EIA)

P.O.R.T.S

Aanderaa (Norway)

Installation:

physical characteristics of the equipment;

requirements concerning installation (e.g. supply voltage and frequency, earthing, power consumption, heat dissipation);

safety information;

Operation:

operational controls, functions and presentation of information; startup and shutdown procedures;

tests, failures, calibration procedures and alarms; interfaces and software description;

of updating procedures; safety information;

Failure Elimination, Maintenance and Service:

required tools and materials for failure elimination and maintenance on board; access for maintenance;

preventive maintenance;

replacement, repair, failure elimination, adjustment; information concerning external service.

### MISCELLANEOUS

In this section additional requirements, which are necessary for the individual navigational equipment but could not be subsumed under another headline should be determined here.

### APPENDIX

### REFERENCES

In this appendix all references used in each individual performance standards should be listed.